# Guide to design, construction, and installation of scientific equipment on NOAA WP-3D aircraft

### Introduction

This guide is intended to supply researchers with the basic criteria for instrument design, installation, and operation aboard NOAA WP-3D aircraft.  It is not yet comprehensive, and will be added to over time.  However, deviating from the principles and guidelines included here may result in substantial delay, or cancellation, of your project during integration.  It is the responsibility of the user to comply with all the AOC requirements, and their personnel will assist in interpreting these guidelines to assure conformity.  Experience shows that early, and frequent, communication with AOC engineers and techs during the instrument design process can often save substantial headaches during instrument integration.

### Design load criteria

All fuselage-mounted equipment must be designed or modified to install within the various instrument racks or be secured to existing attachment points. Fuselage-mounted equipment must be demonstrated to be able to withstand the larger of the ultimate flight load or the crash load, where load = weight x load factor.  Load factors specific to the WP-3D aircraft are given in Table 1.

Table 1.  Equipment design load factors.

| Direction | Flight load factor | Crash load factor |
| --- | --- | --- |
| Forward | - | 10 g |
| Up | 4.2 g | 2 g |
| Down | 7.1 g | 5 g |
| Lateral | 1.9 g | 1.5 g |

Data from “NOAA Stress Analysis Worksheets WP-3D”

(Lockheed Corp.)

In addition, externally-mounted equipment must also be able to withstand aerodynamic loads, calculated based on size and shape for sea-level pressure and a maximum speed of 300 knots. The Data and Development Branch at AOC will assist in the interpretation of these criteria.

Equipment weight limits for fuselage-installed instrumentation varies by location, and maximum acceptable weight and overturning moment data are listed in Appendix 1.

AOC-supplied electronics racks accept 19”-standard rack mount cases, which can be obtained from many different manufacturers, e.g., several styles from Techmar or the DII-series cases from Buckeye Enclosures. All user-supplied rack-mounted equipment should be contained fully within boxes whenever possible, and the structural integrity and method of mounting to the rack will be subject to inspection by AOC personnel.  When equipment is not fully enclosed, the PI may be required to provide an engineering analysis documenting the ability to withstand the relevant loads using the factors listed in Table 1.

Finally, please keep in mind that the aircraft cabin door is 27” wide.

### Materials

The pressurized WP-3D aircraft environment can subject scientific equipment to extremes of vibration, shock, heat, and electrical instability.  Acceptable designs incorporate a demonstrable margin of safety using aircraft-compatible materials without undue weight gain.  Commonly accepted materials for aircraft structural elements include heat-tempered (T5 or higher) aluminum alloys 2024, 6061, and 7075.  If a structural aluminum part is welded, it must then be heat-treated to T5 or higher temper before it will be accepted.  Steel alloy 4130 and stainless steel alloys 301, 302, 304, and 316 are acceptable for structural parts.  Other materials such as copper, brass, plastics, iron, and non-tempered aluminum may not be used for structural members.

Structural welds (inlet fairings, rack corner joins, etc.) must be performed either by an FAA-certified welder, or by a non-certified welder and subsequently approved by an FAA-certified inspector.  In the Boulder area, vendors providing this service include:

| Aviation Welding Technology Inc. | Advanced Aviation Services | Aero Systems Inc. |
| --- | --- | --- |
| 1116 Colorado Avenue | 3760 Wheeling St. Suite #11 | 2580 S. Main St. |
| Longmont, CO 80501 | Denver, CO 80239 | Erie, CO 80516 |
| (303) 776-2810 | (303) 371-7579 | (303) 665-9321 |
| (303) 682-5911 fax | (303) 371-0109 fax | (303) 665-6367 fax |
|  | www.advanced-aviation.com | www.asicolorado.com |

Design input from the welder should be solicited prior to any machining for proper weld preparation and materials selection.

### Fasteners

Electronics racks, seats, and other substantial scientific equipment are secured to the floor and/or sidewall using heavy-duty extruded cargo tracks (Brownlines) and 5/16” stud fasteners; the necessary hardware is provided by AOC.  These fasteners can be easily repositioned in 1” increments fore-aft in the aircraft.

Aircraft-rated fasteners (MS, AN, or NAS standards) such as screws, nuts, bolts, and rivets, must be used for all structural assembly.  All threaded fasteners must be secured by self-locking nuts, self-locking inserts, or safety wire.  Rivets or through-holes accepting a bolt, washer, and locking nut are preferred over blind tapped holes.  Because of taper at the ends, installed structural bolts and screws must project by two complete threads for full rated strength.  Locking steel inserts (Helicoils, Keen-serts, etc.; available from McMaster-Carr) must be used for structural tapped holes in aluminum.

Whenever possible, aircraft-rated fasteners should be used for other, non-structural elements of the equipment as well.  When flat-head, countersunk fasteners are used, the 100° countersink is preferred over the 82° countersink.  These and many other aircraft-rated fasteners can be ordered online from a number of vendors, e.g., Genuine Aircraft Hardware, Aircraft Spruce, and Wicks Aircraft Supply.

Standard vacuum fittings without positive locking designs, such as KF or ISO clamps, must be secured by safety wire or a simple tie-wrap to prevent loosening under vibration.  When properly seated, compression fittings such as Swagelok, Gyrolok, Parker etc. are under tension by design and do not normally require additional locking hardware for safe installation.

Weight

With nearly three tons of scientific instruments installed in the fuselage in 2004, we were at the maximum zero-fuel weight permitted for the NOAA WP-3D aircraft.  It is safe to assume that the aircraft, minus our equipment, will not get lighter over time.  Each PI should make it a priority to continue to work towards minimizing the space and weight of the instrumentation where possible.  Some instruments can shed weight more easily than others, and unfortunately there are no hard and fast guidelines to use in making these decisions.

As an example, ca. 10-lb savings have often been realized by switching from steel to aluminum compressed gas cylinders, and from steel to aluminum regulators, with no detriment to a given measurement.  A ca. 250-lb savings is afforded by automating an instrument that would otherwise have required an operator.  Conversely, a pound inappropriately saved from a structural part has often proved costly in time and resources, so please consult before trimming “excess” material from structural members.

Fuselage instrument weight is a zero-sum game; if one instrument gains, that weight must be reduced on a different instrument elsewhere.  We will work with each PI to evaluate potential instrument weight savings measures, while considering the attendant cost in time and resources.  Please feel free to consult if you are interested in assessing weight reduction steps for your instrumentation.

### Electrical Power

Wiring.  The following wire insulation materials are acceptable for scientific equipment installed in the aircraft:  Kynar, PTFE Teflon and ETFE (Tefzel) Teflon, Poly-X (alkane-imide), Stilan, and TKT (Teflon/Kapton/Teflon).  Unacceptable insulation materials include PVC, Nylon, or Kapton.  Wire size relative to current-carrying capacity is a safety concern and the guidelines in Table 2 may never be exceeded.

Table 2.  Current load limits.*

| ## Wire size AWG | Contact size | Max. current for test in amps | Fuse/breaker (max) |
| --- | --- | --- | --- |
| 24 | 20 | 3 | 5 |
| 20 | 20 | 7.5 | 10 |
| 20 | 16 | 7.5 | 10 |
| 16 | 16 | 13 | 20 |
| 14 | 12 | 17 | 25 |
| 12 | 12 | 23 | 30 |
| 10 | 10 | 33 | 40 |
| 8 | 8 | 46 | 60 |

*Assumes stranded copper wire with Teflon insulation and crimp contactors.

Contact information up to AWG 12 from PV MIL-C-26482 Series II

handbook.  Hermetic seal contactors must be derated; contact

AOC tech branch for details.

All line power must be individually switched and fused appropriate to the load and power cord rating.  A convenient solution is the Tyco Electronics W23-X1A1G-xx series thermal circuit breaker, available online from Newark, Wicks Aircraft, or Allied Electronics.

Workmanship of wiring in user-supplied equipment must be of high quality throughout.  Exposed conductors are not permitted, and the use of appropriate connector covers, backshells, and shrink-wrap is mandatory.  Electrical connectors should be strain-relieved and wires should be neatly bundled whenever possible.  Wiring bundles, both internal and external to the instrumentation enclosure, must be neatly routed, strain-relieved, and protected against chafing.  Cabling specific to the aircraft installation is highly recommended, to minimize excess length and to eliminate strained connections.  User-supplied equipment with substandard wiring will be repaired by the user and reinspected by AOC personnel prior to installation in the aircraft.

Suggested power connectors are the mil-spec PT and PW series circular connectors with backshell and strain relief.  Either BNC or D-sub connectors are recommended for signal lines.  Crimp-style connections are preferred over solder connections.  Circular, BNC, and D-sub connectors, their strain reliefs, backshells, and mating components (as well as crimp, insertion, and removal tools) can be easily obtained from numerous vendors including Newark, Digikey, and Allied.

A dedicated vendor for PT/PW-style connectors is Circular Connectors Inc., 3250 Corte Malpaso, Camarillo, CO 93012; tel. 805-987-8145, fax 805-987-8556; they usually have stock and are very helpful in identifying part numbers for mating connectors and associated strain reliefs, etc.

Heaters.  Good safety practice is to have a passive thermostat in series and in thermal contact with the heated unit.  The thermostat is chosen to open at a certain value above the typical setpoint of the heated device, to limit potential hazards from a shorted or runaway heater controller.  Surface-mounted thermostats can be obtained from Airpax and from distributors of Honeywell temperature sensors.  Heated parts must have adequate shielding to protect personnel from injury, and to prevent ignition of surrounding materials.

Power on the aircraft.  To minimize the risk of fire, scientific power is available only when an AOC crewmember is aboard the aircraft; all instruments must be fully powered down when unattended.  On flight days, power is typically available 3 hours before takeoff for instrument warmup,  and 1 hour postflight for instrument shutdown.  Access with power on non-flight days is usually available during normal working hours.  In the field, to permit AOC crew rest, at least one day a week is designated as a “hard down day” with no aircraft access.

Power can be sourced externally, from either the hangar electrical supply or a diesel-powered ground cart, or internally, from the auxiliary power unit (APU) or from the #2 engine.  Switching between these sources occurs on flight days; before takeoff, power source sequence is typically (ground cart –> APU –> #2 engine), and after landing (#2 engine –> APU –> ground cart).  Each power source is subject to instability or failure (hangar loses power, ground cart runs out of fuel), and glitches or outages in switching over from one source to another are common.  Some services, especially 115V 60Hz single-phase, seem especially unstable during power switchover.

Good instrument design takes these power instabilities into account so that scientific equipment is fault-tolerant, i.e., not disrupted by short-term (<1sec) transients or damaged by longer-term outages.  Experience has shown instrument rack-mount computers to be highly sensitive to switching transients and should not be plugged into the 60Hz bus; laptops tolerate the 60Hz transients well, however.  Most, if not all, standard computer power supplies seem to accept 400Hz without a problem.

The scientific power distribution bus is controlled by the flight crew; in emergencies, the flight crew may disconnect scientific power without warning.  In any case, designing and testing(!) a simple power fault recovery procedure BEFORE installation on the aircraft is highly recommended.  This is especially important for unattended instruments, where this procedure would be carried out by other PIs on board.

### Power distribution

### Scientific electrical power in the cabin can be provided via AOC-standard distribution boxes using NEMA receptacles, each rated at 15 Amps (Figure 1).  Available power is 115V 60Hz from either Bus 1A or 1B (depending on location in the aircraft), 115V 60Hz from Bus 2, 115V 400Hz, 28VDC, and 28VDC weight-on-wheels (WOW).

### Figure 1.  AOC power distribution box (receptacles on opposite face not shown).

### The abundant but finite power available on board the aircraft is shown in Table 3, broken out by service bus. AOC engineers have recommended keeping current load on any single bus below 75% of its maximum rated current to balance the total load on the 29-year-old WP-3D frequency converters; buses loaded past this recommended percentage are shown in red, below.

Table 3.  Total WP-3D scientific power.

| Service | NEMA plug | Total, A | Bus: Imax, A | Used in 2004 | % used |
| --- | --- | --- | --- | --- | --- |
| 115V, 60Hz, 1-Ø | 5-15P | 105 | 1A: 35 1B: 35 2 : 35 | 24.9 21.5 19.9 | 71 61 57 |
| 115V, 400Hz, 1-Ø | L5-15P | 150 | A: 50 B: 50 C: 50 | 31.3 44.5 35.0 | 63 89 70 |
| 115V, 400Hz, 3-Ø | (not spec’ed) | 35 | 35 | 29.8 | 85 |
| 28VDC | 6-15P | 150 | 1A: 50 1B: 50 2A: 50 | 10.9 34.0 8.0 | 22 68 16 |
| 28VDC WOW1 | 6-20P | 50 | 2B: 50 | 1.7 | 3 |

1power is OFF when weight is on landing gear.

The 2004 scientific instrumentation loads in Table 3 suggest that very little additional 60Hz or 400Hz power is available.  At present, the most headroom remains on the 28VDC buses, and researchers are encouraged to use these where possible.  For converting aircraft 28VDC to other voltages, DC-DC converters (e.g., the ComPAC line from Vicor) and DC-AC inverters (e.g., the aircraft-rated 28V inverter from AeroMedix.com, click on “power accessories”; the Samlex 24V inverters available at Don Rowe Inc. also accept 28V DC input) are commercially available.

Standard scientific power at wing stations is limited to 115V 400Hz and 28VDC.  Other power requirements (e.g., 115V 60Hz or 28VDC WOW) may be accommodated at particular wing stations on a limited basis; these requests must be communicated to AOC engineers well in advance of instrument integration.

Additional electrical power guidelines:

•  60Hz inductive loads, such as motors, should be fed from Bus 2.

•  DC motors must be brushless to avoid arcing.  Large motors (≥0.5 HP) must be thermally and electrically protected.

• Use of high-voltage equipment (lasers, CRTs, channeltrons, etc.) in unpressurized external pods requires additional consideration to inhibit arcing at low pressures.  These installations must be designed and certified to safely operate at pressures down to 250 Torr.

•  Grounding must be accomplished by a connection to aircraft power ground; the use of a chassis ground to the airframe is not acceptable.

### Data distribution in flight

Sharing data between different systems can be valuable for several reasons, including monitoring a remote system’s status and for monitoring or recording supplemental data generated external to a given instrument.

Aircraft-instrument communications.  A subset of flight-level data (e.g., lat, lon, alt, winds, ambient pressure, ambient temperature) is available via serial data stream for up to xx (Sean?) systems.  The subset of parameters is selected prior to integration. Currently there is a total of xx (Sean?) GPS time information feeds, in the IRIG-B format, for synchronizing instrument computer clocks to the aircraft-standard time.   Instrument software should be designed to be tolerant of occasional aircraft data signal interruption.

Finally, a limited number of analog outputs from individual instruments can be recorded by the aircraft data station, if requested prior to integration.

Instrument-instrument communications.  An internal Category 5/6 Ethernet LAN is dedicated to instruments running a standardized version of DaqFactory control software.  Selected chemical data are available on this network (e.g., CO, O3, NOy, etc.) as well as aircraft flight-level information.  Instruments in this network typically do not require an additional aircraft data serial feed.

Aircraft-ground communications.  While in flight, selected data are periodically transmitted to the ground via ftp using an on-board satellite phone link.  In addition, a second satellite phone is used for messaging via Xchat, permitting up to three different computers to be in text communication with the ground.  Finally, at present, selected computers aboard the aircraft provide full Internet access.  (Some words about whether instruments with individual gps receivers are allowed?  Antenna issues??)

### Environment

Vibration and shock isolation.  Turboprop aircraft flying in the daytime convective boundary layer are subject to relatively high levels of vibration and, to a lesser extent, shock.  Depending on the application, researchers may choose to isolate individual subassemblies, rackmount boxes, or entire racks.  Wire-rope isolators, for which ultimate load data are available, can be obtained directly from Enidine; this brand is recommended for isolation mounts that need AOC approval.  Non-structural isolators are also available from Aeroflex (click on “isolator products”) or Barry Controls.  Tutorials on isolator selection are available from these websites.  When in doubt, please consult with the AOC techs before selecting a vibration mount design and product.

Temperature.  The aircraft environment can subject instrumentation to extremes of heat and cold, and careful thought regarding temperature management will always pay off.  Overheating has been a chronic and serious issue for many instruments in the fuselage, and pod instrumentation running on the ground on a warm, sunny day is also subject to this problem.  Cabin temperatures can be very location-dependent in the aircraft, with measured electronics rack temperatures exceeding 40° C in previous deployments.

Forced air cooling, using lightweight brushless DC fans – e.g., the Flite II series from Comair-Rotron, available from Newark Electronics - can be very effective (up to a point) and is generally recommended.  AOC is working on the cabin cooling issue, but with scientific equipment dissipating 25kW of energy within the fuselage, positive airflow through instrumentation enclosures is often necessary to keep things running properly in an otherwise hot cabin.

### Hazardous materials

Compressed gas cylinders.  Any compressed gas cylinders to be installed in the aircraft must have a current DOT hydrostatic testing date, and all cylinders must be fitted with pressure regulators.  All cylinders must be provided with a positive forward stop; cylinders larger than 1 liter displacement must be held with a minimum of two aircraft-rated cylinder band clamps.  Cylinder clamps can be obtained from several online sites, some of which are listed below.

www.clampco.com – a wide variety of T-bolt and quick-release clamps.

www.avoxsys.com – 800644, 803272 series cylinder band clamp bracket assemblies (for ca. 30 ft3 and larger cylinders).

www.mhoxygen.com - brackets and bands for smaller cylinders.

Please remember that cylinder clamp tensioning screws must be secured with safety wire or tie-wraps to prevent loosening under vibration. Hose clamps are not acceptable for securing compressed gas bottles.

Use of small-volume aluminum or composite cylinders on the aircraft can provide substantial weight savings over larger aluminum or steel cylinders.  Aluminum medical oxygen cylinders (e.g., Cramer-Decker Inc.) are inexpensive, very clean, and are good for 5 years beyond their latest hydrostatic test date.  Composite-wrapped cylinders from Luxfer Composites are lighter than aluminum and are now quite inexpensive, ca. $150 per cylinder, and easily obtained.  Note that aluminum and carbon-fiber-wrapped aluminum cylinders are DOT-rated for 5 years; fiberglass- and Kevlar-wrapped cylinders are only good for 3 years beyond their latest test date.

Compressed gas cylinders must be handled and stored according to safe practices at all times in the hangar and on the aircraft or ramp.  Please do not move cylinders with regulators attached unless properly restrained by a safety chain on a cylinder cart.  Cylinder caps must be securely fastened before bringing pressurized cylinders up the aircraft stairs.  Wheel chocks are required for utility and cylinder carts, especially those used on the ramp near the aircraft; small rubber chocks can be ordered for $8.00 per set from Worden Safety Products, (800) 424-6257, part number ARC-4HP.

Toxic/flammable gases.  Use of compressed toxic and/or flammable gases requires a secondary containment of cylinder and regulator, such that any leak is vented outside the aircraft.  Gas volumes on board should be minimized, and bottles should not be serviced or refilled inside the aircraft; enclosure design should minimize the need to break the enclosure seals to refill or exchange the cylinder inside.  The user must provide a calculation of the aircraft cabin concentrations should the cylinder contents leak out with a concurrent failure of containment seals.  Assume a cabin volume of 4260 ft3 = 1.21 x 105 liters.  Please also provide the OSHA 8-hour PEL and IDLH limits for each toxic substance for comparison.  Battery-powered alarms (e.g., the portable PacIII from Draeger, available from VWR Scientific) for toxic/flammable gases are in most cases mandatory, and must be provided by the user, along with a tutorial for the AOC personnel on their use.

Radioactive materials.  The use of generally licensed ionizing devices has been permitted on a case-by-case basis.  Prior installations have used 210Po (alpha emitter, tau = 138.4 days) for mass spectrometer ion sources, 63Ni (beta emitter, tau = 100 yrs) for GC detectors, and 85Kr (beta emitter, tau = 10.7 yrs) for DMA columns.  Please contact the AOC lead tech at least four weeks before bringing any radioactive source onto base.

#### Fire suppression/detection in pods.  The AOC is currently reassessing its requirements in this regard, and will advise us as these plans are finalized.  At present, fire suppression in each pod is required.  In 2004 this consisted of a fused pyrotechnic canister with a dry chemical suppressive agent installed in each pod; details of the canister can be found at www.pyrogen.com or contact AOC engineer Jorge Delgado at (813) 828-3310 x3033.  It is probable that more extensive pod fire detection systems will be required in the future, so please stay tuned.

# Sampling inlets and probes

Design and construction of chemical and aerosol sampling inlets and probes, as well as optical ports and radiometer collector housings, should be approved six weeks prior to the beginning of installation by AOC personnel.  For instruments installed in the fuselage, window plate blanks are provided and can be machined to accept a variety of inlet designs.  Most instrument inlets have standardized on either a sideways-facing inlet (using 4130 steel streamline tubing extrusion; part no. SL33-14-4, Wicks Aircraft Supply) or a rearward-facing inlet.  Please contact Tom Ryerson if you think one of these might fit your sampling needs.

### Pumps and exhaust

Continuous exhaust flows may not be vented through window plates upstream of any other instrument sampling locations along the fuselage. For 2006, we are working to provide appropriately-sized PFA Teflon lines connected to separate nadir dump ports in the lower instrument port, forward (LIPF, informally known as the “pump bay”), for continuous exhaust flows.  A sketch showing the location and diameter of these planned exhaust lines can be downloaded from the P-3 2006 Houston web site (password required).  Vent ports for non-continuous, or emergency, exhausts (LN2 dewar venting, containment vessel dumps, etc.) may be acceptable upstream of other sampling inlet locations on a case-by-case basis.

Liquid or condensing gases should never be directed into these exhaust lines.  Further, coalescing filters on the exhaust ports of all oil-sealed pumps are required, to minimize contamination over time of the exhaust tubing installed in the aircraft.

At present, two Busch pumps will be installed in the pump bay.  Requests for additional LIPF pump installations, or availability of any excess pumping capacity, should be coordinated with Andy Neuman, our unofficial pump bay czar.

The ring venturi mounted beneath the LIPF is currently reserved for pumping the LTI in 2006.

### Installation

Installation begins approximately 8 weeks before the transit/science flights begin, subject to aircraft availability. New installations will be generally scheduled for early in the integration period.  Please arrive with ALL necessary sheet metal brackets, hardware, mounting plates, safety covers, compressed gas cylinder mounts, and machining work already finished; AOC does not have a machine shop.  Their sheet metal shop is excellent but is easily swamped by multiple last-minute requests for their services.  Please register all your requests for sheet-metal support on the WP-3D Installation Worksheet well in advance of arriving in Tampa.

### Equipment inspection and acceptance

AOC will need to better define how we can get things officially accepted – investigators have commented on a confusing process so far, with any number of AOC personnel seeming to have a veto on any given installation, right up until the flight crew walks through at the last moment before test flights commence.  Is there a better way we could jump through 99% of the hurdles well in advance?

Miscellaneous information

•  Tool control:  Flight-critical wiring, control cables, hydraulic lines, and actuators run overhead, along the sidewalls, and underneath the floor inside the P-3 fuselage.  As a result, tool control is strictly implemented on the P-3 as a flight safety measure.  If you plan to use any tools during installation or operation on the aircraft, please provide a complete list to SED before bringing them on board, and check the tools against this list before departing each day.  “Tool control” also extends to any small parts – nuts, washers, hardware, jewelry, etc. – taken on board the aircraft.  If you drop something on the floor, you are expected to pick it up before continuing with other tasks.  If you can’t immediately locate a dropped item, please call it to the attention of the AOC techs, who will be happy to assist in locating it with you.

#### •  Hygiene & footwear:  Work in and around the aircraft can entail long, hot hours in a poorly ventilated space in close proximity to other scientists, engineers, and air crew.  Good attention to personal hygiene will be expected and very much appreciated by all your colleagues.  While safety shoes are not required, they are highly recommended; people wearing open-toed sandals or shoes are never allowed on, or near, the aircraft.

•  Documentation:  It is always good practice to bring your own copies of MSDS, structural analyses, documentation, wiring schemes, etc., relevant to your installation.  While AOC may have their own copy, substantial down time can sometimes be saved if you have a backup copy in your own possession when needed.

### •  Weather.  Please anticipate all-weather operation with regard to chemical sampling inlets and preflight/postflight operations.  Hangar access has occasionally been arranged, but plan on sheltering necessary equipment from rain (and occasionally snow) to ensure a smooth, complication-free field deployment.

### Appendix 1.  P-3 fuselage equipment load limitations

| Station | Max weight (lbs) |  |  | Mounting location (flight station) |  |  | CG location (inches from nose) |  |  | Overturning moment (in-lbs) |  |  | Vib. # |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Inboard | Outboard | Total | Forward | Aft | Delta | # Flight station | W.L. | Height | Max inboard | Max outboard | Total |  |
| FD | 312.5 | 312.5 | 625.0 | 300.1 | 319.1 | 19.0 | 311.85 | 155.5 | 32 | 10000 | 10000 | 20000 | 6 |
| C3X | 350.0 | 350.0 | 700.0 | 429.0 | 465.0 | 27.0 | 453.9 | 151.5 | 28.0 | 9800 | 9800 | 19600 | 8 |
| Sta. 2 | 250.0 | 250.0 | 500.0 | 459.0 | 486.0 | 27.0 | 477.19 | 150.3 | 26.8 | 6700 | 6700 | 13400 | 8 |
| Sta. 3 | 272.0 | 385.0 | 657.0 | 526.0 | 553.0 | 27.0 | 541.2 (±) | 154.4 | 30.9 | 8405 | 11396 | 19801 | 8 |
|  |  |  |  |  |  |  | 543.0 (0) | 153.1 | 29.6 |  |  |  |  |
| Sta. 4 | 332.0 | 291.0 | 623.0 | 604.9 | 628.1 | 23.2 | 619.9 (±) | 151.8 | 28.3 | 9396 | 8264 | 17660 | 8 |
|  |  |  |  |  |  |  | 620.0 (0) | 151.9 | 28.4 |  |  |  |  |
| N1/N2 | 504.0 | 430.0 | 934.0 | 675.0 | 700.0 | 25.0 | 685.8 (±) | 161.1 | 37.6 | 18950 | 13287 | 32237 | 8(1) |
|  |  |  |  |  |  |  | 686.1 (0) | 154.4 | 30.9 |  |  |  |  |
| Sta. 5 | 350.0 | 350.0 | 700.0 | 710.6 | 737.6 | 27.0 | 727.1 | 151.5 | 28.0 | 9800 | 9800 | 19600 | 8 |
| Sta. 6 | 262.5 | 262.5 | 525.0 | 776.1 | 799.9 | 23.8 | 792.7 | 153.5 | 30.0 | 7875 | 7875 | 15750 | 6 |
| Sta. 7 | 262.5 | 262.5 | 525.0 | 841.1 | 864.9 | 23.8 | 857.6 | 153.5 | 30.0 | 7875 | 7875 | 15750 | 6 |
| Sta. 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Dual pass. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| J cab |  |  |  |  |  |  |  |  |  |  |  |  |  |

Refs: 	(1):  NOAA Equipment and Interface Load Criteria – CALAC 1-14-77

(2):  NOAA Equipment Interface Load Criteria – CALAC 3-14-76

From a document signed by J. Zysko, 27 April 81

### Appendix 2.  P-3 fuselage station locations